1575 Commands and SCPI issues/questions

ESE and ESR commands are simular. ESR says it will be cleared by the query, but that doesn't appear to happen.

Page 3 (2.3) – Execution error messages are are listed in XXX what is XXX, is the :SYST:ERR? Codes on page 90 (3.0)

Page 94 (4.4-4.5) – Two instances of :STB? when it should be *STB? unless there is a secret :STB command.

What does and doesn't STB clear?

Does QUE clear “:STAT:QUES:EVEN?”?

If each command should be checked for “ERROR”, how long do we need to wait for it, as it could significantly slow down the speed of sending commands?

PS and PT can be report to be stable, but the dual stable is not set immediately.

How do I know when non-remote called self test is called?

“Inside a loop or some other timer mechanism, check for errors and for
status. Errors could be overpressure or loss of control, which typically
indicates a leak. Should either of these conditions exist, the system will
automatically put itself into VENT mode and the program should have code
to handle this condition appropriately.”
It doesn't appear to vent on over-pressure, why not GTG?